## Parts List
MicroLab Parts List
NOTE 1: We link to places that have the parts that we used. They aren't picked for best price or universal availability
NOTE 2: Saving money - check the price of a part in multiples, it may be cheaper than buying just one
Section | Part ID | Count | Name | Primary Link | Secondary Link | Description / Notes
CONTROL UNIT (CU)
Control Unit (CU) | CU-RPI | 1 | Raspberry Pi 3 Model B Board | Amazon link | Amazon Le Potato link2 | Primary link is to a Raspberry 3b+'s which is the most extensively tested. Second link is for devlopers, we have experimental support for Le Potato by Libre Computer (link2) although as of v0.6 we do not yet have a production image with display drivers installed.
Control Unit (CU) | CU-SCR | 1 | Touchscreen for Raspberry Pi | Amazon link |  | Touchscreen for Raspberry Pi (CU-RPI). Minimum recommended screen resolution is 640 x 480.
Control Unit (CU) | CU-SD32 | 1 | Micro SD card, 32GB | Amazon link |  | 32GB is the minimum. If you make a new image on a larger card, you can use Pi Shrink to remove empty space from the disk image. If you shrink to the minimum, you may need to re-expand or your partition will run out of space.
Control Unit (CU) | CU-UPC | 1 | Micro USB Cable, right angle | Amazon Link |  | Power cable for Raspberry Pi (CU-RPI). The right angle (90 degree) connector will fit better in the Control Unit case. You will need a cable that fits your Pi, we linked to a Micro USB (although newer board usually use USB C).
Control Unit (CU) | CU-UNO | 1 | Arduino UNO | Ali Express link (kit) - be sure to pick option with UNO board | Amazon link | Links to kit will include Uno, CNC shield, motor drivers modules andle USB A/B cable.
Control Unit (CU) | CU-CNC | 1 | Arduino CNC shield |  | Amazon Link (Kit)
Control Unit (CU) | CU-SMD | 3 | Stepper Motor Drivers [For CU-CNC] |  | Amazon Link | Link 1 includes UNO and 4988 but 4988KIR may be of lesser quality
Control Unit (CU) | CU-A2B | 1 | USB A/B cable for Raspberry Pi to Arduino
Control Unit (CU) | CU-RLY | 1 | 4-Channel Relay Board | Amazon link
Control Unit (CU) | CU-WCL | 1 | Wire Connector (lever), 2-in-6-out | Ali Express link | Amazon link - assortment | Called the 12V Wire Connector. The 2-in-6-out wire connector is used for 12V wiring. The 65-pack in the second link is enough to make multiple MicroLabs, smaller packs are avaiable but will be a higher cost per connector.
Control Unit (CU) | CU-JMP | 5 | Breadboard Jumper Wires, male-female | Amazon link |  | Small wires to connect the pins on the Pi to other components.
Control Unit (CU) | CU-RLY | 1 | 4 Channel Relay Board, 5V | Amazon link
Control Unit (CU) | CU-POW | 1 | Power Supply, 12V 5A | Amazon link |  | This a 12v 5A power brick. If you’ve got one from an old laptop with a 5.5 x 2.1 mm plug, it may work just fine. Powers the MicroLab.
Control Unit (CU) | CU-CBC | 1 | Control Case (Printed) | MicroLab link | Github Link | The MicroLab index page goes into more detail about the 3D printed parts.
Control Unit (CU) | CU-CBL | 1 | Control Lid (Printed) | MicroLab link | Github Link | The MicroLab index page goes into more detail about the 3D printed parts.
Control Unit (CU) | CU-CCP | 1 | Converter case (Printed) | MicroLab link | Github Link | The MicroLab index page goes into more detail about the 3D printed parts.
Control Unit (CU) | CU-RCP | 1 | Relay case (Printed) | MicroLab link | Github Link | The MicroLab index page goes into more detail about the 3D printed parts.
Control Unit (CU) | CU-PLG | 2 | Barrel plug connectors, 5.5mm x 2.1mm | Amazon link |  | Connector with wires, 1 female and 1 male.
Control Unit (CU) | CU-USB | 2 | Mountable USB cables | Ali Express Link | Amazon Link | Straight connectors will work best, shorter length (25cm) is recommended.
Control Unit (CU) | CU-SDC | 1 | 12V to 5V Converter, stepdown | Amazon link | Ali Express Link | Stepdown converter with barrel connector and USB ports.
Control Unit (CU) | CU-RAH | 1 | Right angle header row | Amazon link
Control Unit (CU) | CU-WCS | 1 | Wire Cutter and Stripper | Amazon link
Control Unit (CU) | CU-FCT | 1 | Flush cutters | Amazon link
Control Unit (CU) | CU-NNP | 1 | Needle-nose pilers | Amazon link
Shared Parts
Shared Parts | SP-CAB | 15ft roll | 4-wire cable, 18 gauge | Amazon link |  | [Control Unit] [Pumps Box] 10 feet will be used to make cables to connect the Control Unit and the Pumps unit. Some of the rest will be used to wire components within the Control Unit. You can also use wires from ethernet cables.
Shared Parts | SP-RBW | 40ft roll | 2-wire cable, red and Black , 18 gauge | Amazon link |  | Optional and we won't use all of it, you may want to get a smaller amount if you can find it cheaper. Or just use ethernet cable.
Shared Parts | SP-SWW | 15ft roll | Split Wire Wrap, half-inch | Amazon link |  | [Control Unit] [Pumps Box]
Shared Parts | SP-PC12 | 2 | 12-pin connector, panel-mount screw terminal | Ali Express link |  | [Control Unit] [Pumps Box] connector has three pieces: plug, mount plate and socket
Shared Parts | SP-PC8 | 2 | 8-pin connector, panel-mount screw terminal |  |  | [Control Unit] [Pumps Box] connector has three pieces: plug, mount plate and socket
Shared Parts | SP-PC2 | 1 | 2-pin connector, panel-mount screw terminal |  |  | [Control Unit] connector has three pieces: plug, mount plate and socket
Shared Parts | SP-M3S-12 | 23 | 12mm M3 screws with nuts (Optional) | Amazon link - assortment |  | [Reactor Unit] 20 for reactor stand (Optional)
Shared Parts | SP-M3S-16 | 28 | 16mm M3 screws with nuts |  |  | [Control Unit] [Pumps] 20 for Pumps Box, 8 Control Unit
Shared Parts | SP-M3S-20 | 2 | 20mm M3 screws with nuts |  |  | [Control Unit] [Reactor] 2 for Control Unit
Shared Parts | SP-M3S-30 | 5 | 30mm M3 screws with nuts |  |  | Pumps Box 4
Shared Parts | SP-SMC | 6 | Stepper motor cables | included w/ peristaltic pumps |  | The 3 stepper motors come with their own cable, but will need to cut in half and both ends stripped. The linked ones are ready to use.
Shared Parts | SP-ETN | 1 roll | Electrical tape, narrow | Amazon link |  | 1" is most common, 1/2" should be narrow enough
TEMPERATURE CONTROL UNIT
Temp Control (TC) | TC-PMP | 2 | Circulating pump | Amazon link |  | To run both heating and cooling in the same recipe without hardware reconfiguration between recipe steps, 2 pumps are required. NOTE: Operational limits to pumps in terms of capacity and temperature have not been determined. Larger pumps may be desired for optimal operation, especially at high temperatures. Self-priming pumps are HIGHLY recommended. NOTE: These will actually be installed in the Pumps Unit.
Temp Control (TC) | TC-CTB | 2 | Copper tubing 10’ ¼" OD x 3/16" ID | Amazon link for raw copper tubing | Amazon link for aluminum cooler | Copper tubing should fit snugly within the silicone tubing. This 10’ length should be enough to make 1 dense or 2 loose heat exchanger coils. Pre-formed coils used for chilling in beer brewing are more expensive, but easier. These are also typically too large for even wide-mouth mason jars, so another container may be needed. We are also looking into pre made aluminum coolers although aluminum is almost 50% less conductive than copper!
Temp Control (TC) | TC-STB | 1 | Silicone tubing 8mm OD x 5mm ID | Amazon link |  | Silicone tubing should fit snugly around the copper tubing
Temp Control (TC) | TC-TAP | 1 | Tape for the insulation
Temp Control (TC) | TC-TCT | 1 | Copper tube cutter tool | Amazon link
Temp Control (TC) | TC-CTJ | 1 | Copper tube jig |  |  | Currently we are using the Reactor Core manifold as a jig, other options include, a cylinder with a smaller diameter than the jar or a medium-to-large dowel/tube
Temp Control (TC) | TC-LIQ | 1 | Liquid, propylene glycol, water, or other |  |  | We have been using propylene glycol which support temperature ranges as high and as low as the MicroLab supports.
REACTOR UNIT
Reactor Unit | RU-SRM | 1 | Stir rod motor | Amazon link |  | Amazon link This link is to a 100 RPM 12v motor. You can also buy a faster motor and use a PWM controller to adjust the speed manually.
Reactor Unit | RU-STR | 1 | Stir rod | Amazon link |  | Mixing paddle needs to be coated / PTFE. Stainless steel may react. Paddle should be small enough to fit easily in the 6oz mason jar
Reactor Unit | RU-SRC | 1 | Stir rod coupler 4mm-to-5mm | Amazon link
Reactor Unit | RU-PTFE | 1 | PTFE Sheet | Amazon Link
Reactor Unit | RU-RJF | 2 | Jar Flanges (Printed) | MicroLab link | Github Link
Reactor Unit | RU-COR | 1 | 250ml GL45 Borosilicate Glass Bottles |  | Amazon link | NOTE: New Reactor Core uses a GL45 250ml bottle instead!
Reactor Unit | RU-OUT | 1 | 32oz wide mouth mason jar for outer jacket | Amazon link
Reactor Unit | RU-SDH | 1 | Reactor stand - H (Printed) | MicroLab link | Github Link | H shaped piece for the reactor stand
Reactor Unit | RU-SDA | 1 | Reactor stand - A (Printed) | MicroLab link | Github Link | A shaped piece for the reactor stand
Reactor Unit | RU-SDC | 1 | Reactor stand - ¢ (Printed) | MicroLab link | Github Link | ¢ shaped piece for the reactor stand
Reactor Unit | RU-SDK | 1 | Reactor stand - K (Printed) | MicroLab link | Github Link | K shaped piece for the reactor stand
Reactor Unit | RU-TMP | 1 | Thermistor | Ali Express link |  | Used to monitor the temperature in the reaction chamber. This is a link to a USB DS18B20
Reactor Unit | RU-PPM | 3 | Peristaltic pumps with stepper motors | Amazon link | Alibaba Link | The 3 stepper motors should come with their own cable, but will need to cut in half and split between the Control Unit and the pumps box. NOTE: Those links are likely to be sold out. There are other stepper-driven pumps, but they will not interface seamlessly with the printed parts. We’re also looking for DIY/printable designs.
Reactor Unit | RU-SYR | 3 | Syringes | Amazon link |  | 20 or 30mL are standard sizes. You may want more of these or different sizes.
Reactor Unit | RU-PUB | 1 | Pumps box (Printed) | MicroLab link | Github Link | Case to put pumps and motors in
Reactor Unit | RU-PUL | 1 | Pumps lid (Printed) | MicroLab link | Github Link | Lid for pumps case
Reactor Unit | RU-SUP | 1 | Wire coat hanger |  |  | Used to support the syringes

## Sheet2 (hidden)
pic label | ID code | Name | count | Primary Link | Second link | Model / Part # (mcmaster carr) | Description / notes | Price Aug/Sep 2022 | From
01 | MC-WNT | wire nuts | 2 | link |  | 8904T3 | Lever connectors are better, orange wire nuts (video) will also work | 7.1 | McMasterCarr
02 | SP-TB4 | 4mm x 6mm vinyl tubing (1/8") |  | link |  |  | or 1/8 inch | 5.49 | amazon
03 | RX-PMP | circulating pump | 2 | link | link2 |  | Mini pump (first link) may not withstand being sumberged in 100C water. The 5A pump in link 2 is almost certainly overkill, but should withstand high heat and is self-priming. Power supply connector here. Power supply here. We may want smaller pumps that can run at fewer amps (screw terminals seem to max out at 12v 5A, which is only half of what two of the larger pumps require. | 25.9 | lightobject | 47.98 | amazon
04 | RX-TMP | thermistor | 1 | link |  |  |  | 7.36 | aliexpress
05 | RX-TMPD | USB dongle for thermistor | 1 | link |  |  |  |  | same link as thermistor
06 | SP-SLD | syringe slides | 4 | link |  |  |  | 9.99 | amazon | right item?
06A | SP-DRV | M5 threaded drive rod | 2 | link |  |  | appropedia calls for 200mm, thread pitch not specified | 8.99 | amazon
07 | RX-MXM | mixing motor | 1
08 | RX-PDL | mixing paddle | 1 | link1 | link2 |  | Mixing paddle needs to be coated. Stainless steel may react. | 25.99 | amazon | not available | amazon
09 | RX-TB5 | tubing 5-6mm x 8mm (1/4") | 1 | link |  |  | pump outlet | 5.99 | amazon
09B | RX-TB8 | tubing 8mm x 12mm |  | link' |  |  | pump inlet | 13.99 | amazon
09C | RX-RED | Reducing fitting |  |  |  |  | might be needed
10 | RX-TEE | barbed tee fittings for 6mm tubing | 2 | link |  | 2429K55 |  | 3.82 | McMaster Carr
11 | MC-SCR | touch screen | 1 | link |  |  | screen is to a raspberry pi 4 | 36.99 | amazon
12 | MC-RLY | relay hat | 1 | link |  |  |  | 10.91 | aliexpress
13 | MC-RPI | raspberry pi | 1 | link |  |  | Primary link is to a 4 - 3b+'s seem hard to come by | not available | amazon | ~$12.00
14 | MC-SD32 | micro SD card, 32GB | 1 | link |  |  | 32GB | 13.25 | amazon
15 | MC-MDR | motor drivers for CNC hat | 4 | link |  |  | Link includes UNO and 4988 | 64.08 | aliexpress
16 | MC-CNC | arduino CNC hat | 1 | link |  |  |  |  | same link as motor drivers for CNC hat
17 | MC-UNO | arduino UNO | 1 | link |  |  |  | 26.79 | amazon
18 | SP-TAB | plunger holder tab (Linear Actuator V0 7) | 2 | link |  |  |  |  | print instuctions not a thing to buy
19 | SP-BDY | body holder (Linear Actuator V0 5) | 2 | link |  |  |  |  | print instuctions not a thing to buy
20 | SP-PGR | plunger holder base (Linear Actuator V0 7) | 2 | link |  |  |  |  | print instuctions not a thing to buy
21 | SP-END | motor end (Linear Actuator V0 1) | 2 | link |  |  |  |  | print instuctions not a thing to buy
22 | SP-CAR | carriage (Linear Actuator V0 4) | 2 | link |  |  |  |  | print instuctions not a thing to buy
23 | SP-IDL | idler end (Linear Actuator V0 2) | 2 | link |  |  |  |  | print instuctions not a thing to buy
24 | RX-LID | reactor manifold assembly | 1 | printed |  |  |  |  | no print instructions
25 | SP-MTR | stepper motors | 2 | link | link | 6627T91 |  | 13.7 | aliexpress | 13.99 | amazon
26 | SP-M310 | M3x10mm screw |  | link |  |  | Amazon assortment contains all needed sizes and wrench | not available | amazon | ~$25
| SP-M320 | M3x20mm screw
| SP-M340 | M3x40mm screw
| SP-M3N | M3 nut
26A | SP-M5B | 625Z Ball Bearing (5mm x 16mm x 5mm) | 2 | link |  | 6153K15 |  | 11.98 | amazon
27 | RX-COR | 4/6 oz mason jar for core chamber | 1 | link | 6oz link |  |  | 7.63 | amazon | 19.99 | amazon
28 | MC-A2B | USB A/B cable for rPi to Arduino | 1 | link |  |  |  |  | same link as motor drivers for CNC hat
29 | SP-SYR | syringes | 2 | link |  |  | 25mL luer slip tip in video, 20 or 30mL are standard sizes. luer lock will also work | 1.18 | CA Vet Supply
30 |  | hex wrench for syringe pump, M3 | 1
31 | RX-OUT | 1qt wide mouth mason jar for outer jacket | 1 | link |  |  |  | 9.35 | amazon
| MC-GFCI | 15 Amp GFCI Outlet Adapter |  | link |  |  | Safety precaution since relays have 120v mains | 16.99 | amazon
| MC-WIRE | 18 gauge wire for power supply | 1 | link |  |  |  | 8.99 | amazon
336.46
[8 empty] | 37 | item n/a from amazon estimated for substitutes
[8 empty] | 3.9 | shipping | lightobject
[9 empty] | shipping | McMaster Carr- won't say without progressing order
[6 empty] | based on my zip code |  | 4 | shipping | aliexpress
[8 empty] | 9.85 | shipping | CA Vet Supply | cheapest option
[8 empty] | free if slow shipping | shipping | amazon
391.21